## Slide 1: Data analysis and filtering
Professor Fabrice Pierron
University of Southampton

## Slide 2: Summary
Performance evaluation
  Resolution, spatial resolution
Strain calculation
  Smoothing, filtering, reconstruction
Performance reporting
Mechanical data extraction

## Slide 3: Performance evaluation (1/10)
Metrology
  Science of measurement
  Recommendations (vocabulary, definitions, good practice) available in GUM
http://www.bipm.org/en/publications/guides/gum.html
Guide to the expression of uncertainty in measurement

## Slide 4: Performance evaluation (2/10)
Two types of errors: Random and Systematic
Accurate but not Precise
Large Random Error
Small Systematic Error
Precise but not Accurate
Small Random Error
Large Systematic Error
Words used interchangeably
Bias = Systematic error
Noise = Variance error = Random error
Image from Wikipedia
The new view is to categorize as Type A and Type B Errors
Copyright 2013 Sandia Corporation
Notes: This is because a true bias error can be corrected, which leaves only a random error (along with any error involved in the correction).

## Slide 5: Performance evaluation (3/10)
A newer approach categorizes errors into two types: Type A and Type B
Chart: Parameter Value
| 76.379900000000006 | 76.381600000000006 | 76.383300000000006 | 76.385000000000005 | 76.386700000000005 | 76.388400000000004 | 76.390100000000004 | 76.391800000000003 | 76.393500000000003 | 76.395200000000003 | 76.396900000000002 | 76.398600000000002 | 76.400300000000001 | 76.402000000000001 | 76.403700000000001 | 76.4054 | 76.4071 | 76.408799999999999 | 76.410499999999999 | 76.412199999999999 | 76.414000000000001 | 76.415700000000001 | 76.417400000000001 | 76.4191 | 76.4208 | 76.422499999999999 | 76.424199999999999 | 76.425899999999999 | 76.427599999999998 | 76.429299999999998 | 76.430999999999997 | 76.432699999999997 | 76.434399999999997 | 76.436099999999996 | 76.437799999999996 | 76.439499999999995 | 76.441199999999995 | 76.442899999999995 | 76.444599999999994 | 76.446299999999994 | 76.447999999999993 | 76.449700000000007 | 76.451400000000007 | 76.453100000000006 | 76.454899999999995 | 76.456599999999995 | 76.458299999999994 | 76.459999999999994 | 76.461699999999993 | 76.463399999999993
Counts - Plot 0 | 0.50024 | 0.71565 | 1.0072 | 1.3944 | 1.8991 | 2.5444 | 3.3536 | 4.3483 | 5.5463 | 6.9595 | 8.5907 | 10.4319 | 12.4618 | 14.6446 | 16.93 | 19.2539 | 21.5408 | 23.7076 | 25.6681 | 27.339 | 28.6453 | 29.5261 | 29.9391 | 29.8645 | 29.3057 | 28.2899 | 26.8654 | 25.0978 | 23.0654 | 20.8529 | 18.5462 | 16.2265 | 13.9661 | 11.8252 | 9.8497 | 8.0709 | 6.5058 | 5.1589 | 4.0244 | 3.0884 | 2.3315 | 1.7315 | 1.265 | 0.90916 | 0.6428 | 0.44708 | 0.3059 | 0.2059 | 0.13634 | 0.08881
Type A – Evaluated via statistical methods
Repeated measurements
Statistical distributions
  Normal, Log-normal, etc.
Copyright 2013 Sandia Corporation
Notes: Both error types are propagated via the same sum-squared methods.

## Slide 6: Performance evaluation (4/10)
A newer approach categorizes errors into two types: Type A and Type B
Mathematical modeling of the experiment taking into account all error sources is a valid and approved method of estimating uncertainty. (Section 3.4.1)
, …, )
The measurand Y is made up of X other input quantities. The function may be so complicated that it cannot be written down (Section 4.1.2).
Type B – Evaluated by other means
Modeling approaches
Assumed probability distribution
Experimental expertise
Copyright 2013 Sandia Corporation
Notes: Both error types are propagated via the same sum-squared methods.

## Slide 7: Performance evaluation (5/10)
Resolution
  Smallest value of the measurand (displacement, strain, slope etc…) that can be detected above noise
  Different from sensitivity: smallest value of measurand associated with a change of signal (typically related to CCD sensitivity and digitization here)
  Resolution is above sensitivity here because of camera noise

## Slide 8: Performance evaluation (6/10)
Resolution
  Evaluate by mean and standard deviation of measurement on a still object in real testing conditions
    Error on geometrical model can be accounted for by performing rigid body mouvements
    Bias can be evaluated by a change of magnification (see DIC lecture)
  Always report these values when communicating results

## Slide 9: Performance evaluation (7/10)
Spatial resolution
  Smallest distance between to INDEPENDENT measurements
  In DIC: spatial resolution in displacement is related to the subset size (not the shift!)
  In the grid method, spatial resolution is related to the grid pitch
  In interferometry, spatial resolution is the pixel
only 2 independent data points

## Slide 10: Performance evaluation (8/10)
Spatial resolution
  Previous definition not convenient (DIC/grid complex non-linear filters)
  Frequency cut-off approach
Cut-off frequency

## Slide 11: Performance evaluation (9/10)
Practical implementation
  Just before performing your test, record at least two images of the still specimen (grid or DIC)
  Calculate displacement using the DIC or grid software
  Calculate the standard deviation of this displacement map. This can be used to estimate the resolution in displacement
  Check that the mean is lower than the standard deviation
  Report the resolution in displacement with your data

## Slide 12: Performance evaluation (10/10)
Noise
  Camera dependent !!
  Example: Shimadzu HPV-1 (Ultra high speed camera)

## Slide 13: Spatial resolution (1/9)
1000 x 1000 images, 8 bit
41
31
21

## Slide 14: Spatial resolution (2/9)
Analytical displacement field
Synthetically deformed images
DIC and comparison with reference displacement of 1 in amplitude
2D DIC, affine shape function, SSD criterion, bicubic spline interpolation (MatchID)

## Slide 15: Spatial resolution (3/9)
Effect of subset size: 21 (step 50%)
Amplitude: 0.996 - Error: 0.4%

## Slide 16: Spatial resolution (4/9)
Effect of subset size: 31 (step 50%)
Amplitude: 0.972 – Error : 2.8%

## Slide 17: Spatial resolution (5/9)
Effect of subset size: 41 (step 50%)
Amplitude: 0.942 – Error : 5.8%

## Slide 18: Spatial resolution (6/9)
Spatial resolution for subset size 41 (step 50%)
Chart: Reconstruction error (%)
5 | 7.8
8 | 16.3
11 | 29.1
13 | 38.8
15 | 48.3
17 | 59.9
Resolution: 0,012 (standard deviation of displacement noise map, max of components)
5% error

## Slide 19: Spatial resolution (7/9)
Effect of shape function (subset size 41, step 50%) - quadratic
Chart: Reconstruction error (%)
5 | -0.7
8 | -0.5
11 | 0.4
13 | 1.7
15 | 3.4
17 | 6.1
Resolution: 0.041 (standard deviation of displacement noise map, max of components)
5% error

## Slide 20: Spatial resolution (8/9)
Effect of step size
Large step size
Small step size

## Slide 21: Spatial resolution (9/9)
Effect of step size
Better sampling but same spatial resolution!

## Slide 22: Strain calculation (1/20)
From displacements to strains
  Strains (small deformation, small strain)
  Simplest way to compute strains: finite differences
L

## Slide 23: Strain calculation (2/20)
From displacements to strains
  Differentiation amplifies noise
x
x
x
x
x
s
e
E1
E2E1
E

## Slide 24: Strain calculation (3/20)
From displacements to strains
  One thing to remember: strain is a relative displacement over a gauge length
  To improve resolution:
    Increase resolution on DL
    Increase the gauge length L
  Some basic figures:
    If displacement resolution is 0.01 pixel, then resolution on DL is 0.02 pixels. If the strain is computed between two DIC subsets set apart by say 20 pixels, the raw strain resolution will be 2.10-3
    To improve it, smoothing is necessary

## Slide 25: Strain calculation (4/20)
From displacements to strains
  Finite differences
x
y

## Slide 26: Strain calculation (5/20)
From displacements to strains
  Finite differences

## Slide 27: Strain calculation (6/20)
Smoothing
  Low pass filter
FFT
Modulus of FFT
Spatial frequency
Signal
Noise
Signal+ noise
Noise + signal
FFT
Spatial frequency cut-off: regularization parameter

## Slide 28: Strain calculation (7/20)
Mean filtering
  Replace central value by mean of neighbours (kernel n). Improve resolution by
  Regularization parameter: kernel
1 | 1 | 1
1 | 1 | 1
1 | 1 | 1

## Slide 29: Strain calculation (8/20)
Gaussian smoothing
  Replace central pixel value by weighted average of neighbouring pixels
  Removes outliers, low pass filter, tuning parameter: kernel

## Slide 30: Strain calculation (9/20)
Gaussian smoothing
  Kernel of 8
Problem with the edges

## Slide 31: Strain calculation (10/20)
Finite element smoothing
  Use piecewise functions to reconstruct the displacements
  Identify nodal values using least-square approach
FE mesh
Size 10 pixels

## Slide 32: Strain calculation (11/20)
Global polynomial smoothing
  Find the polynomial coefficients that fit the data best in the least-square sense
  Strain by finite difference of smoothed data (or analytically from polynomial)
  Tuning parameter: polynomial degree
  Very efficient low pass filter
  Instabilities towards the edges
  Strain at one point depends on displacements all over the specimen!

## Slide 33: Strain calculation (12/20)
Global polynomial smoothing
  Degree 8

## Slide 34: Strain calculation (13/20)
Local polynomial smoothing
Least square fit

## Slide 35: Strain calculation (14/20)
Local polynomial smoothing
  Virtual Strain Gauge size (VSG)
Virtual strain gauge size
VSG = [ (SW-1) x ST] + SS
Depends on both strain window and subset size
SW = Strain window (in data points)
ST = Step size (pixels)
VSG = 19 pixels
Here:
Subset (SS) = 7 ; Step (ST) = 3, Strain window (SW) = 5

## Slide 36: Strain calculation (15/20)
Spatial resolution vs resolution
  Smoothing improves resolution and degrades spatial resolution
  Smoothing spreads noise spatially (blurs information)
  Strain resolution meaningless without the spatial resolution (‘gauge length’).
  IMPORTANT: regularization parameter (degree, kernel, VSG etc…) more important than smoothing method!

## Slide 37: Strain calculation (16/20)
Effect of smoothing
Subset = 29
Step = 3
Strain Filter = 5
Virtual Gage = 1.6 mm
Subset = 71
Step = 5
Strain Filter = 5
Virtual Gage = 2.7 mm
Subset = 71
Step = 5
Strain Filter = 15
Virtual Gage = 8.2 mm
Subset = 29
Step = 8
Strain Filter = 15
Virtual Gage = 13.1 mm
Subset = 91
Step = 23
Strain Filter = 15
Virtual Gage = 37 mm
Copyright 2013 Sandia Corporation

## Slide 38: Strain calculation (17/20)
Example
Copyright 2013 Sandia Corporation

## Slide 39: Strain calculation (18/20)
Virtual Strain Gauge = 50 pixels
Spatial resolution

## Slide 40: Strain calculation (19/20)
Subset = 21 pixels

## Slide 41: Strain calculation (20/20)
Spatial resolution in strain
  Standard procedure feasible for all techniques still missing
  Idea: spatial frequency approach
    Define a cutting frequency in the Fourier space
    Work underway
  In any case, always precise method and tuning parameter when reporting results

## Slide 42: Performance evaluation
Always evaluate performances before test
Minimal procedure
  Take two sets of still images (better: out of plane movement)
  Compute displacement
    Compute standard deviation: displacement resolution
  Compute strains
    Compute standard deviation: strain resolution
  Report key parameters
    Subset for DIC, grid size (Grid Method), camera type, camera spatial resolution etc.
  Report resolution (in the future, add spatial resolution)

## Slide 43: Reporting
Always report performances with data

## Slide 44
Extracting mechanical data from strain fields

## Slide 45: Statement of the problem
Problem to solve
Force
Geometry
+
Analytical solution
Virtual Fields Method (see this afternoon)
Extract E and n ?

## Slide 46: Speckle deformation simulation (1/2)
DIC is a rather complex and non-linear process; the results of depend on many parameters, including the speckle pattern
Any realistic UQ on the mechanical parameters requires to take into account the measurement process: numerical speckle image deformation tool
Start with FE generated displacement map and a speckle pattern (real or artificially generated)

## Slide 47: Speckle deformation simulation (2/2)
Principle
New pixel intensity: average of
Doing this for all pixels leads to a deformed image
DIC leads to simulated displacements affected by the correlation


## Slide 48: Displacement error
Results
y
Bias: low spatial frequency error
1.5 pixel displacement
1.5 sine period
Vertical displacement in pixels

## Slide 49: Strain error
y
VSG=151 pixels
SR=211 pixels
Larger error
(larger gradients)

## Slide 50: Strain error with noise on images
y
Much larger than systematic error
VSG=151 pixels

## Slide 51: UQ on parameters / deformation (1/5)
Illustration
  Synthetic strain map containing the effect of, say, local stiffness variations in the material (the 1 to 4 figures)
integration
DIC
Rossi M., Lava P., Pierron F., Debruyne D., Sasso M., International Journal of Solids and Structures, Strain, 2015.

## Slide 52: UQ on parameters / deformation (2/5)
Global reconstruction error
  Wrong criterion
Reconstruction error (RMS): 0.0035
Reconstruction error (RMS): 0.0017

## Slide 53: UQ on parameters / deformation (3/5)
When performing material identification
  Error on material parameters are relevant, not so much error on strains
  Illustration
Rossi M., Pierron F., International Journal of Solids and Structures, 2012.
FE model, glass/epoxy UD composite
Field of view in red

## Slide 54: UQ on parameters / deformation (4/5)
Shear strain maps obtained by image deformation plus different levels of smoothing
VSG=5 pixels
(no smoothing)
VSG=60 pixels
(local polynomial)
Which one is ‘the best’?

## Slide 55: UQ on parameters / deformation (5/5)
Because reference stiffness components are known, an error function can be built-up (RMS error)
10
(grey level)
Lower error
Larger error
30
Same error
Same error
150
Lower error
2 x error

## Slide 56: Smoothing vs reconstruction error (1/3)
Systematic error coming from low pass filtering effect of measurements
  Illustration: ‘Arcan’ test on foam (tension/shear on rectangular specimen)
Wang P., Pierron F., Thomsen O.T., Experimental Mechanics, 2013
Wang P., Rossi M., Lava P., Pierron F., Thomsen O.T., Strain, 2015
FE strains
Simulated DIC + smoothing

## Slide 57: Smoothing vs reconstruction error (2/3)
The total error is the sum of
  Systematic error (or bias, or reconstruction error), from the low pass filtering effect of the measurements. It increases with VGS
  Random error arising from noise, decreases with VSG
DIC parameters can be optimized using synthetic measurements
Systematic error obtained by using noise free images (reference and synthetically deformed)
Random error obtained by performing many identification with different copies of noise, taking the mean and subtracting the systematic error

## Slide 58: Smoothing vs reconstruction error (3/3)

## Slide 59: Experimental validation – 1/3
Comparison of strain maps
FE/DIC simulation with noise and reference stiffness
Experimental

## Slide 60: Experimental validation –2/3
  Mean, average over 20 repetitions (results in MPa)
Chart: 
Reference | 143.4
Simulator | 146.1
Specimen 1 | 149.3
Specimen 2 | 147.6
Specimen 3 | 150.1
Chart: 
Reference | 63.4
Simulator | 66
Specimen 1 | 68.1
Specimen 2 | 70.8
Specimen 3 | 69.1
Chart: 
Reference | 26
Simulator | 27.9
Specimen 1 | 27.9
Specimen 2 | 29.6
Specimen 3 | 29.9
Chart: 
Reference | 30.1
Simulator | 30.9
Specimen 1 | 32.9
Specimen 2 | 31.1
Specimen 3 | 32.4

## Slide 61: Experimental validation – 3/3
  Coefficients of variation, average over 20 repetitions (results in MPa)
Chart: 
Simulator | 0.18419
Specimen 1 | 0.23374
Specimen 2 | 0.25916
Specimen 3 | 0.28268
Chart: 
Simulator | 1.4581
Specimen 1 | 2.0056
Specimen 2 | 2.0116
Specimen 3 | 2.4545
Chart: 
Simulator | 0.73477
Specimen 1 | 1.0203
Specimen 2 | 0.98662
Specimen 3 | 1.2079
Chart: 
Simulator | 0.39482
Specimen 1 | 0.76292
Specimen 2 | 0.73955
Specimen3 | 0.85802

## Slide 62: Conclusion
DIC technique: complex non-linear filter
Post-processing: another layer of non-linear filtering
Important to systematically evaluate and report performances (in papers, reports etc)
Post-processing depends on objectives
  Strain evaluation
  Stiffness / constitutive model identification
  Acceleration measurements
  ...