## Slide 1: Metrology and uncertainty quantification in DIC
Professor Fabrice Pierron
University of Southampton

## Slide 2: Introduction (1/3)
Digital Image Correlation (DIC) is becoming a common tool in experimental mechanics
However, relatively new: commercial systems around 10 years old
Long measurement chain, highly non-linear
Many parameters: speckle, camera (noise), lens (distortion), lighting (heating, saturation …), subset size, step size, virtual strain gauge size
Influence of the choice of these parameters? Uncertainty estimation?

## Slide 3: Introduction (2/3)
The three important components of 2D-DIC
Matching
Interpolation
Shape Function

## Slide 4: Introduction (3/3)
DIC: easy and fast (or is it??)
Objectives
  Review each of these three components
  Understand the influence of some of the choices made there on results
  Understand how other choices are conditioned by these three components
  Provide understanding to make these choices in an informed way

## Slide 5: Matching (1/11)
Finding the same pattern (one thing just like the other)

## Slide 6: Matching (2/11)
Using interpolation and the shape function a correlation criterion is used to find a subset match
Subset to find
Region of Interest (ROI)
Reference Image (F)
Matched Image (G)
χ – is the function to minimize
F – is the reference image
G – is the deformed image
i – is the pixel in the subset
Etc.
How do you pick your correlation criterion?
Sum Squared Difference (SSD)
Normalized Sum Squared Difference (NSSD)
Zero Normalized Sum Squared Difference (ZNSSD)
Copyright 2013 Sandia Corporation
Notes: There are a number of other functions that can be used: Cross-correlation, etc.

## Slide 7: Matching (3/11)
The minimization function finds the motion and deformation of a subset simultaneously
How to select …. Speed vs. Image quality
F : reference image
G : deformed image
Summation over all pixels in subset
The parameter vector s relates coordinates in the reference image to coordinates in the deformed image.
A simpler notation for rigid-body-motion only.
Copyright 2013 Sandia Corporation

## Slide 8: Matching (4/11)
Selecting your correlation criterion: Speed versus image quality
Copyright 2013 Sandia Corporation

## Slide 9: Matching (5/11)
Images always have different intensities during the test. Particularly in 3D!
Full Contrast
Reduced contrast
Reduced and shifted contrast
Copyright 2013 Sandia Corporation

## Slide 10: Matching (6/11)
SSD does not compensate for illumination differences
SSD Full Contrast
∆ = 0.94
SSD Reduced & Shifted
∆ = 0.60
SSD Reduced Contrast
∆ = 0.31
χ – is the function to minimize
F – is the reference image
G – is the deformed image
i – is the pixel in the subset
∆ = 1 minus the minimum
Copyright 2013 Sandia Corporation

## Slide 11: Matching (7/11)
NSSD is able to compensate for an offset in the image contrast
NSSD Full Contrast
∆ = 0.97
NSSD Reduced & Shifted
∆ = 0.91
NSSD Reduced Contrast
∆ = 0.97
χ – is the function to minimize
F – is the reference image
G – is the deformed image
i – is the pixel in the subset
∆ = 1 minus the minimum
Copyright 2013 Sandia Corporation

## Slide 12: Matching (8/11)
ZNSSD performs equally on both shifted and compressed contrast
χ – is the function to minimize
F – is the reference image
G – is the deformed image
i – is the pixel in the subset
They are all the same – regardless of image lighting changes.
∆ = 1 minus the minimum
Copyright 2013 Sandia Corporation
ZNSSD Reduced & Shifted
∆ = 0.88
ZNSSD Reduced Contrast
∆ = 0.88
ZNSSD Full Contrast
∆ = 0.88

## Slide 13: Matching (9/11)
ZNSSD performs much better on a wide range of image contrast variations
SSD
0.5029
0.4971
ZNSSD
Copyright 2013 Sandia Corporation
0.855
0.20
0.62
0.407
Displacement in pixels
0.5036
0.4961
0.5036
0.4961
0.5062
0.4918
SSD

## Slide 14: Matching (10/11)
The reported matching criterion (sigma) also shows that ZNSSD is better than SSD
SSD
ZNSSD
Copyright 2013 Sandia Corporation
Notes: Sigma is the match quality reported by the DIC software. It is representative of the matching error.

## Slide 15: Matching (11/11)
How do you choose your correlation algorithm?
  Lighting quality
  Solution speed – this has been ignored in the previous study
  ZNSSD is numerically more complex
Reference Image
Deformed Image
SSD or NSSD or ZNSSD
Which correlation algorithm?
NSSD or ZNSSD
ZNSSD
?
Copyright 2013 Sandia Corporation

## Slide 16: Interpolation (1/9)
Interpolation – The information between the pixels (‘Reading between the lines’)

## Slide 17: Interpolation (2/9)
Interpolation is a technique to find the data between sampled points.
Chart: Pixel Location
Chart: Pixel Location
Copyright 2013 Sandia Corporation

## Slide 18: Interpolation (3/9)
The images are interpolated to be able to calculate a subpixel match
Image interpolated by a factor of 10x
210 x 210 pixels
21 x 21 pixels
Copyright 2013 Sandia Corporation

## Slide 19: Interpolation (4/9)
You have a number of interpolation options: but not all of them are good!
Linear
Cubic Polynomial
Cubic Spline
Fourier Transform
Optimized filter (4-Tap, etc.)
What constitutes a good interpolator?
Matches the value at pixel locations
Minimizes the phase error
Minimizes the amplitude error
i.e. Minimizes the error in the matching of shifted images!
Copyright 2013 Sandia Corporation

## Slide 20: Interpolation (5/9)
The differences between the images are subtle but important
FFT vs FIR (~4-Tap)
σ = 0.61
Original Image
FFT vs Linear
σ = 3.55
FFT vs Cubic Spline
σ = 1.43
σ – Standard Deviation of the grey values.
σ = 0 would be a perfect match
Copyright 2013 Sandia Corporation
Notes: The FFT has “more support” therefore is closer to the perfect sinc interpolation. All other interpolation functions are truncations and modifications of this.

## Slide 21: Interpolation (6/9)
Bias and random error: illustration
  Horizontal uniform elongation of a square FOV, displacement of 0 at the left, 1 at the right
Exact
Simulated DIC
Difference

## Slide 22: Interpolation (7/9)
A major error source in DIC uncertainty quantification!
Schreier et al., Systematic errors in digital image correlation caused by intensity interpolation. Opt. Eng. 39, 2915-2921 (2000)
1-D Translation
Copyright 2013 Sandia Corporation

## Slide 23: Interpolation (8/9)
Legacy interpolants caused phase and amplitude errors in the results
Copyright 2013 Sandia Corporation

## Slide 24: Interpolation (9/9)
Newer interpolants are no longer the largest error source in the DIC UQ chain
Copyright 2013 Sandia Corporation

## Slide 25: Shape functions (1/5)
Subset shape functions (‘Constraining your solution’)

## Slide 26: Shape functions (2/5)
Like finite element methods, the shape function matches the subset displacement
The affine Subset:
Translations (u, v)
Rotation
Scaling
Shearing
In some other codes higher ordered subsets exist. What are the advantages/disadvantages?
Copyright 2013 Sandia Corporation

## Slide 27: Shape functions (3/5)
Undersized subsets can lead to problems matching the shape or deformation
A linear shape function on a circle
Pascal Lava – KU Leuven

## Slide 28: Shape functions (4/5)
Large subsets lead to lower spatial resolution

## Slide 29: Shape functions (5/5)
Another approach is “global” DIC, where the surface is meshed and solved with an FE formulation.
Tomičevć, Z., et al. (2013). "Mechanics-aided digital image correlation." The Journal of Strain Analysis for Engineering Design 48(5): 330-343.
Surface is meshed with elements.
Nodes are constrained to improve noise floor
X-FEM can deal with cracks.
Higher-order elements are OK.
A “regularization” length is still needed.
A priori information: dedicated shape functions

## Slide 30: Uncertainty in DIC
The matching error is composed of a bias and variance term
Interpolation Bias
Systematic
Bias
Noise Bias
Measurement Variance
Random
Variance
21000
-21000
0
Gradients =
=
Noise 5%
Noise 1.5%
Copyright 2013 Sandia Corporation
interpolation

## Slide 31: Interpolation bias (1/10)
Effect of aliasing
Average speckle size
~0.3 mm
~1 mm

## Slide 32: Interpolation bias (2/10)
Out of plane mouvement
  Induces artificial ‘strain’ only in xx and yy components
  A good way to check for aliasing problems
stage
Correlate the two set of images to evaluate aliasing error

## Slide 33: Interpolation bias (3/10)
Results
  5 mm out of plane translation
  Interpolation: bi-cubic polynomial

## Slide 34: Interpolation bias (4/10)
Results
  5 mm out of plane translation
  Interpolation: bi-cubic polynomial

## Slide 35: Interpolation bias (5/10)
Results
  5 mm out of plane translation
  Interpolation: bi-cubic polynomial

## Slide 36: Interpolation bias (6/10)
5 mm translation
10 mm translation

## Slide 37: Interpolation bias (7/10)
Effect of pre-filtering
No prefiltering
Gaussian 5x5 pixels
prefiltering

## Slide 38: Interpolation bias (8/10)
Prefiltering blurs sharp edges
  Easier interpolation
  Unfortunately, increases noise
Interpolation Bias
Systematic
Bias
Noise Bias
Measurement Variance
Random
Variance

## Slide 39: Interpolation bias (9/10)
Example
Strongly aliased
x
Linear interpolation
Gaussian prefiltering (7x7)
No prefiltering

## Slide 40: Interpolation bias (10/10)
Effect of interpolation
Bicubic splines
Bicubic polynomials

## Slide 41: Noise bias (1/3)
Interpolation Bias
Systematic
Bias
Noise Bias
?
Measurement Variance
Random
Variance

## Slide 42: Noise bias (2/3)
Same speckle pattern, two different amounts of noise (2 and 8 grey levels)
Pattern rigidly translated by steps of 0.1 pixel
Mean displacement calculated for each step
Plot mean against translation step

## Slide 43: Noise bias (3/3)
Results
  Bi-cubic polynomial, no pre-filtering
Displacement in pixels
Shift in pixels
8 grey levels
2 grey levels

## Slide 44: Random variance (1/3)
Stationary pattern with increasing amounts of noise
  Calculate standard deviation of displacement

## Slide 45: Random variance (2/3)
Effect of image averaging
  Take many images and average them for both deformed and undeformed
  Illustration: 50 images of stationary pattern
  DIC between image 0 and image 1
Standard deviation of
u: 0.0064 pixel
v: 0.0066 pixel

## Slide 46: Random variance (3/3)
Effect of image averaging
  DIC between average of images 0 to 25 and average of images 25 to 50
Standard deviation of
u: 0.0019 pixel
v: 0.0017 pixel
Reduced by 3!

## Slide 47: A practical problems (1/4)
Perpendicularity of camera wrt specimen
specimen
CCD sensor
Out of plane displacement

## Slide 48: A practical problem (2/4)
Perpendicularity of camera
  Motorized stage, control angle
  Perpendicularity attained when displacement is uniform

## Slide 49: A practical problem (3/4)
Example: vertical translation of 10 mm
Displacement non-perpendicular
Displacement perpendicular

## Slide 50: A practical problem (4/4)
Example: vertical translation of 10 mm

## Slide 51: Other cases (1/5)
Stereo-image correlation
  Matching issues same as 2D DIC
  Extra error sources: 3D calibration process
  Some issues become obsolete: perpendicularity, out of plane movements
  UQ is more complex because of calibration
  Still under development

## Slide 52: Other cases (2/5)
Digital Volume Correlation
  Extension of DIC to volume data
Different sources
  X-ray CT
  Confocal microscopy
  Optical Coherence Tomography (OCT)
  MRI
  Sheet illumination for transparent bodies
  …

## Slide 53: Other cases (3/5)
Examples: X-ray CT of bone*
*: Gillard et al., JMBBM, 2014

## Slide 54: Other cases (4/5)
Examples: OCT of porcine cornea*
*: J. Fu, P.D. Ruiz, F. Pierron (PhD of J. Fu), 2014
Inflation test

## Slide 55: Other cases (5/5)
Digital Volume Correlation
  Same issues as 2D matching
    Interpolation, aliasing, noise etc.
  Stationary and rigid body motion tests to evaluate performances
  More difficult to induce ‘strain’ by change of magnification (under way)
  Uncontrolled ‘speckle’ pattern: related to material
  Imaging system induces ‘distortions’ (refraction in OCT, ring artefacts in X-ray CT etc.)

## Slide 56: Conclusion
DIC is a complex measurement technique
Many choices have to be made, which impact the results
No magic recipe: need for understanding the different issues
Qualitative vs quantitative DIC: a step change
Need for training
More in the data processing lecture

## Slide 57: Some useful resources
International DIC society: http://idics.org
Guide of ‘good practice’ (iDICs)
  https://doi.org/10.32720/idics/gpg.ed1
Training
MatchID DIC course
  Six successful 5 days sessions in 2014-2018
  Next edition: 1-7 July 2019, Ghent, Belgium
  https://www.matchid.eu/en/dic-course

## Slide 58: Acknowledgements
Dr Pascal Lava, MatchID N.V., Belgium
Dr Philip Reu, Sandia National Laboratory, USA